## Slide 1: Mechanical Ventilator
Professor Yasser Mostafa Kadah
EE 471

## Slide 2: Mechanical Ventilator
A ventilator is a life-sustaining device that supports or replaces spontaneous breathing of the patient

## Slide 3: Spontaneous Breathing Cycle
A breath is one cycle of positive flow (inspiration) and negative flow (expiration) defined in terms of the flow versus time curve

## Slide 4: Gas Exchange: Diffusion
Gas is exchanged between alveoli and blood by diffusion
Speed of diffusion is determined by:
  Partial pressure gradient (+)
  Thickness of barrier (-)
In cases when fluid accumulates inside alveoli, hyperoxic (more O2) is needed to compensate for increased barrier thickness

## Slide 5: Equations of Motion: Electrical Analogy
Mechanics Electricity
Pressure  Voltage
Flow  Current
Volume  Charge
Resistance  Resistance Compliance  Capacitance
Valve  Switch

## Slide 6: Practical Effects of Compliance and Resistance
A: Filling of normal lung unit
B: Low-compliance unit: fills quickly but with less air
C: Increased resistance: fills slowly
Inspiration
Expiration

## Slide 7: Functional Block Diagram

## Slide 8: Oxygen Sources
Medical gases supply
Oxygen cylinder
Oxygen concentrator

## Slide 9: Air/Oxygen Blender
Gas mixer allow the user to vary oxygen concentration of inspiratory gas or fractional inspiratory oxygen (FIO2) between 21% and 100% by volume
  Mechanical gas mixers
  Electronically-controlled gas mixer integrated in ventilator
Gas mixers are usually responsible for ensuring that breathing gas to be supplied is prepared and delivered in required quantity and rate

## Slide 10: Pressure/Flow Generator
Pressure/flow generator is responsible for delivering mixed gas prepared by gas mixer according to selected ventilation parameters
Flow generator is a controlled valve whose output provides defined gas flow with output pressure is not specified
Pressure generator behaves similar to compressor, whose output provides defined pressure with unspecified gas flow

## Slide 11: Breathing System
Breathing system forms interface between patient and ventilator
Clinical ventilators are usually connected to patient via inspiratory and expiratory hoses (dual-hose circuit)
Gas flow delivered through inspiratory port passes through breathing gas humidifier before entering patient’s lungs
  To make it adapted to climatic conditions in patient’s lungs
After inspiratory phase, patient exhales through expiratory valve
  Breathing systems of ventilators are referred to as non-rebreathing circuits

## Slide 12: Gas Humidifier
Humidifiers are used to warm and humidify inspiratory gas.
  Dry and cool supply gas dry out the patient’s airways with risk of causing irreversible damage to ciliated epithelium
Active gas humidifiers
  Located in the inspiratory limb and use electrical energy to heat a water bath
  When cold, dry gas passes over the water surface it absorbs water molecules and is thus warmed and humidified
  Examples: pass-over humidifiers and bubble-through humidifiers

## Slide 13: Gas Humidifier
Passive breathing gas humidifiers or heat and moisture exchangers (HMEs)
  Designed to buffer significant fraction of moisture and heat expired by patient
  Retained moisture is then used to condition inspired gas passing through HME during next inspiration

## Slide 14: Expiratory (Exhalation) Valve
Allows expired air to move out of breathing circuit
If valve is not opened completely during expiration, positive end-expiratory pressure (PEEP) is created in lungs
  PEEP is therapeutically important as it increases gas exchange surface of lungs
  Adequate PEEP can also prevent collapse of individual alveolar areas
If expiratory valve is controlled during inspiratory phase, it can compensate for undesired pressure rises in breathing system
  Caused, for example, by patient coughing

## Slide 15: Patient Interface
Noninvasive Ventilation (NIV) vs. Intubation

## Slide 16: Infection Control: HEPA Filtration
Filtration is the passage of a liquid or gas through a filter with pores small enough to retain microbes
  Membrane filters composed of cellulose esters are commonly used to filter out bacteria, viruses
Microbes can be removed from air by high-efficiency particulate air (HEPA) filters
  Remove almost all microorganisms larger than about 0.3 m in diameter

## Slide 17: Trigger, Limit, Cycle, and Baseline Variables
Trigger variable is one that is measured and used to start inspiration
Limit variable is one that can reach and maintain a preset value before inspiration ends (i.e., does not end respiration)
Cycle variable is one that is measured and used to end inspiration
Baseline variable is the parameter controlled during expiration
  Pressure control is most practical and used in all modern ventilators

## Slide 18: Modes of Mechanical Ventilator
“Mode” is defined as a predetermined pattern of interaction between a ventilator and a patient
  There are over 100 names for modes of ventilation on commercially available mechanical ventilators
  Neither the manufacturing community nor the medical community adhere to a standard taxonomy for modes
Volume (VC)
Pressure (PC)
CMV
IMV
CSV
Set-point
Adaptive
Servo
Intelligent
Ventilator Mode
+
+
=
Breath Control Variable
Breath Sequence
Targeting Scheme

## Slide 19: Modes of Mechanical Ventilator

## Slide 20: Beginning of Inspiration: Trigger Variable
Time triggering
Patient triggering
  Pressure
  Flow
  Volume

## Slide 21: During Inspiration: Control and Limit Variables
Ventilator can control one variable
  Pressure, volume, flow, or time
Limit variable is maximum value variable can attain during inspiration
  Does not end inspiratory phase
  Ventilator can have multiple limit variables
Control
Control

## Slide 22: Termination of Inspiration: Cycle Variable
Volume-Cycled Ventilation
Time-Cycled Ventilation
Flow-Cycled Ventilation
Pressure-Cycled Ventilation

## Slide 23: Targeting Scheme

## Slide 24: Example Ventilator Modes
CPAP
Spontaneous
Pressure-Support (PS)
Continuous Mandatory Ventilation (CMV)
Synchronized Intermittent Mandatory Ventilation (SIMV)

## Slide 25: CPAP, PEEP and BiPAP
Continuous Positive Airway Pressure (CPAP)
Positive End-Expiratory Pressure (PEEP or EPAP)
BiLevel Positive Airway Pressure (BiPAP)
BiPAP
CPAP/PEEP + Spontaneous
PEEP + CMV
CPAP/PEEP + IMV

## Slide 26: Mechanical Ventilator Functional Diagram
Sensors
  Pressure
  Flow
  Oxygen
  Carbon dioxide
Actuators
  Valves
  Air blower

## Slide 27: Pressure Sensor
Piezoresistive Sensor
  Solid-state device whose electrical resistance depends on strain of membrane (and hence, pressure)
  Can be designed to be absolute or differential
  Can be selected to have very fast response time
Resistive strain gauge

## Slide 28: Flow Sensor: Pneumotachograph
Sensing change in pressure across fixed resistance through which gas flow is laminar to calculate flow rate
  Analogous to measuring current through resistance in electric circuits
  Bidirectional operation

## Slide 29: Flow Sensor: Variable Orifice Flowmeter
Pressure measurement across flow-dependent resistance (flap)
  As gas flow increases, flap opens more, reducing resistance to gas flow
  One-to-one correspondence between flow rate and pressure drop
  Bidirectional operation

## Slide 30: Flow Sensor: Hot-Wire Anemometer
Resistive wire is placed in flow path and heated using electric current
  Resistance increases linearly with temperature
Flow cools down wire and changes its resistance
  Resistance change can be measured and used to estimate flow rate
Fast and accurate but direction insensitive

## Slide 31: Flow Rate Sensor: Vane Anemometer
Used to measure exhaled tidal volume in breathing circuit near exhalation valve

## Slide 32: Oxygen Sensor: Fuel Cell or Galvanic Cell
Diffusion of oxygen starts chemical reactions
Flow of current depends on uptake of oxygen at the anode (Faraday’s first law of electrolysis)
  Voltage output is proportional to oxygen partial pressure
Response time (~20-30 s) and depleted by continuous exposure to oxygen because of exhaustion of the cell (less than 1 year lifespan)

## Slide 33: Oxygen Sensor: Paramegnetic
Oxygen is paramagnetic while other gases are diamagnetic
Magnetic field causes oxygen molecules to be attracted and agitated
  Pressure difference proportional to oxygen partial pressure
Very accurate, highly sensitive and fast response
  Allows continuous monitoring
Long lifespan with no service requirements

## Slide 34: CO2 Gas Sensor (Capnography)
Infrared absorption spectroscopy
  Based on physical principle that gases absorb infrared radiation at characteristic frequencies
  Level of absorption depends on concentration of molecules according to Beer-Lambert law: Concentration = Absorption constant × ln(I0/I )
Measures End-Tidal Carbon Dioxide (EtCO2)
Io
I

## Slide 35: Proportional Solenoid Valves
Idea: constriction causes significant pressure drop
Control of pressure/flow of gases
  Inspiratory valve (control of ventilation mode delivery)
  Expiratory valve (Control of PEEP)

## Slide 36: Solenoid Pinch Valves
No contact with air
  No need for sterilization after use
  Particular application: expiratory valve
Proportional control of flow

## Slide 37: Check Valves
Ensure unidirectional flow in both inspiratory and expiratory limps
Mechanical operation

## Slide 38: Adjustable Pressure Limiting (APL) Valve
Safety of airway pressure during positive pressure ventilation to avoid barotrauma
Usually a mechanical device but can be implemented as a solenoid valve

## Slide 39: Pressure Reduction Valve
Essential for oxygen cylinder interface to ventilators
  Reduction of pressure to 4.1 bar
Mechanical operation

## Slide 40: Air Blower
Common technology: Centrifugal Blowers
  Ideal to maintain continual gas transfer and for creating high pressure from small volumes of air
Speed controllable
Silent operation

## Slide 41: Electronic Gas Blender
Set values: FIO2 and tidal volume
Proportional valves for both air and oxygen
  Flowrate in each is proportional to its control signal
Measure pressure and temperature for each to compensate for volume changes
  Gas law: PV= nRT
Solve 2 equations in 2 unknowns to determine flowrates of oxygen and air
Use feedback from oxygen sensor to correct
Target tidal volume = (Flowrate(O2)+Flowrate(air)) x Inspiratory time
Control Signal 1
Control Signal 2
Isolated Drivers
Air
O2
Oxygen Sensor
Feedback

## Slide 42: Typical Ventilator User Interface
Mode
IPAP/EPAP
Rate
FIO2
Tidal volume
Pressure, flow, volume
Alarm messages

## Slide 43: Ventilator Safety Alarm System
Level 1 events include life-threatening situations
  Loss of input power or ventilator malfunction (e.g., excessive or no flow of gas to the patient)
  Alarms in this category should be mandatory (not subject to operator choice), redundant (multiple sensors and circuits), and non-canceling (alarm continues to be activated, even if event is corrected, and must be reset manually)
Level 2 events can lead to life-threatening situations if not corrected in a timely fashion
  Such things as blender failure, high or low airway pressure, auto-triggering, and partial patient circuit occlusion, and may include suspicious ventilator settings such as an inspiratory-to-expiratory timing (I:E) ratio greater than 1:1
  Alarms for level 2 may not be redundant and may be self-canceling (i.e., alarm inactivated if event ceases to occur)
Level 3 events affect patient– ventilator interface and may influence the level of support
  Such events as changes in patient compliance and resistance, changes in patient respiratory drive, and auto-PEEP
Level 4 events reflect the patient condition alone rather than ventilator function
  Usually detected by stand-alone monitors, such as oximeters, cardiac monitors, and blood-gas analyzers
  Some ventilators are able to incorporate such readings (e.g., capnography) in their displays and alarm systems

## Slide 44: Example Practical Alarm System

## Slide 45: Alarm User Interface
Alarm Setting
Alarm Fatigue
“While device alarms can improve medical care by arming clinicians with valuable information about their patient’s health, most are false or nonactionable, often resulting in a sensory clutter that overwhelms clinicians and caregivers, desensitizing them to the alerts”, AAMI

## Slide 46: Advanced Issues: Breathing Circuit Model
Weak link in patient–ventilator system is patient breathing circuit
  Expensive ventilator with state-of-the-art computer control is connected to patient (priceless) with $2 piece of plastic tubing whose design has not changed appreciably in 20 years
  Resistance and compliance of delivery circuit make flow control and volume delivery more difficult
Patient
Patient
Breathing Circuit
Breathing Circuit
Mechanics Electricity
Pressure  Voltage
Flow  Current
Volume  Charge
Resistance  Resistance Compliance  Capacitance

## Slide 47: Advanced Issues: Leakage Compensation
Leaks are inherent to patient interfaces and circuits in NIV
  Monitoring and compensating for leaks associated with NIV is critical for maintaining the appropriate trigger and cycle levels for the ventilator
  This is critical to assessing mask seal efficacy and preventing overtightening of the mask that may result in pressure sores
Uncompensated leaks generate auto-triggering, prolonged cycling, ineffective inspiratory efforts, and poor pressurization capacity of ventilator

## Slide 48: Recommended Readings
J.M. Cairo, Pilbeam's Mechanical Ventilation: Physiological and Clinical Applications, 5th Ed., Elsevier, 2012.
Martin J. Tobin, Principles And Practice of Mechanical Ventilation, 3rd Ed., McGraw Hill Medical, 2012.