## Slide 1: Nuclear Medicine
Prof. Yasser Mostafa Kadah
EE 472

## Slide 2: Nuclear Medicine Basics
Nuclear medicine scans a very small amount (e.g., nanograms) of radioactive material called a radiotracer is injected intravenously into the patient
  Agent then accumulates in specific organs in the body
  How much, how rapidly and where this uptake occurs are factors which can determine whether tissue is healthy or diseased
Three different modalities under the general umbrella of nuclear medicine
  Plain scintigraphy
  Single photon emission computed tomography (SPECT)
  Positron emission tomography (PET)

## Slide 3: Nuclear Medicine Basics
Relative to most other imaging modalities, nuclear medicine scans (in particular planar scintigraphy and SPECT) are characterized as having:
  Poor SNR
  Low spatial resolution (~5–10 mm)
  Slow image acquisition
  Extremely high sensitivity
  Very high specificity (no natural radioactivity from the body)
  Intrinsic functional characteristics of the information content of the images

## Slide 4: Operation of a Nuclear Medicine Gamma Camera

## Slide 5: Radioactivity
Radioactive isotope is one which undergoes a spontaneous change in composition of the nucleus, termed disintegration
Radioactivity is measured in units of curies (Ci), where one curie equals 3.73x1010 disintegrations per second

## Slide 6: Half-Life Time
Half-life (½): time required for the radioactivity to drop to one-half of its value
Biological half-life of the radiotracer (½,bio) (how long the radiotracer remains in the body) must also be considered
Effective half-life (½,eff) is a combination of the two half-lives

## Slide 7: Radioactivity Example:
Two patients undergo nuclear medicine scans. One receives a dose of radiotracer A and the other radioatracer B. The half-life of A is 6 hours and of B is 24 hours. If the administered dose of radiotracer A is three times that of radiotracer B, and the biological half-lives of A and B are 6 and 12 hours respectively, at what time is the radioactivity in the body the same for the two patients?
Obtain values for decay constants A, B
Obtain values for required time t
Solution:

## Slide 8: Properties of Ideal Radiotracers
Radioactive half-life that is short enough to produce significant radioactivity without requiring a very large initial dose, but not so short that there is significant decay before the required post-injection delay to allow the radiotracer to clear the blood and distribute in the relevant organs
Decay should be via emission of a mono-energetic -ray without emission of alpha- or beta-particles.
  Mono-energetic -ray allows discrimination between Compton scattered and unscattered -rays, thereby improving image contrast. Alpha- or beta-particles are completely absorbed within tissue, therefore increasing the radioactive dose without giving any useful image information

## Slide 9: Properties of Ideal Radiotracers
Energy of -ray should be greater than ~100 keV so that reasonable proportion of -rays which are emitted deep within tissue have sufficient energy to travel through body and reach detector
Energy of the -ray should be less than ~200 keV so that the rays do not penetrate the thin lead septa in the collimator (which is analogous to the anti-scatter grid in X-ray imaging)
Radiotracer should have high uptake in organ of interest and relatively low non-specific uptake in rest of the body
  These two factors lower required dose for the patient and increase image contrast, respectively

## Slide 10: Example Radiotracers
Most widely used radiotracer is 99mTc which is involved in over 90% of planar scintigraphy and SPECT studies
  It exists in metastable state (one with reasonably long half-life)
  Formed from 99Mo
  Energy of emitted -ray is 140 keV

## Slide 11: Technetium Generator
Tc generator is delivered to nuclear medicine department at the beginning of the week, returned at the end of the week
  Very convenient method for producing 99mTc

## Slide 12: Radioactivity of 99mTc vs Time
Generator that is not milked
Generator that is milked every 24 hours

## Slide 13: Gamma Camera

## Slide 14: Collimator
Spatial Resolution (R)
Efficiency (g)
Parallel Collimator

## Slide 15: Detector Scintillation Crystal
In the gamma camera, the detector is a large single crystal of thallium-activated sodium iodide, NaI(Tl), 40–50 cm diameter
When a -ray strikes the crystal, it loses energy through photoelectric and Compton interactions which result in a population of excited electronic states within the crystal
De-excitation of these states occurs ~230 nanoseconds later via emission of photons with a wavelength of 415 nm (visible blue light), corresponding to a photon energy of ~4 eV
A very important characteristic of scintillators such as NaI(Tl) is that the amount of light (the number of photons) produced is directly proportional to the energy of the incident -ray

## Slide 16: Photomultiplier Tube (PMT)

## Slide 17: Photomultiplier Tube (PMT)
Produces between 105 and 106 electrons for each initial photoelectron, creating an amplified current at PMT output
Since each PMT has a diameter of 2–3 cm, and the NaI(Tl) crystal is much larger in size, a number of PMTs are closely coupled to the scintillation crystal.
  Most efficient packing geometry is hexagonally-close-packed
  Arrays of 61, 75 or 91 PMTs are typically used, with a thin optical coupling layer used to interface the surface of each PMT with the scintillation crystal
  Each PMT should ideally have an identical energy response. If this is not the case, then artifacts are produced in the image. For planar nuclear medicine scans, a variation in uniformity of up to 10% can be tolerated whereas only values less than 1% can be tolerated in SPECT

## Slide 18: Anger Position Network
Whenever a scintillation event occurs in NaI(Tl) crystal PMT closest to scintillation event produces largest output current
  If this were only method of signal localization, then spatial resolution would be no finer than dimensions of the PMT, i.e., several cm
Adjacent PMTs produce smaller output currents, with amount of light detected being approximately inversely proportional to distance between scintillation events and particular PMT
  By comparing magnitudes of currents from all PMTs, location of scintillation within crystal can be much better estimated

## Slide 19: Anger Position Network
In modern Anger position networks, interpolation between outputs from adjacent PMTs is used to improve resolution of scintillation detection
Linear Interpolation Example

## Slide 20: Pulse Height Analyzer (PHA)
Pulse Height Analyzer discriminates recorded events into two scenarios
  𝛾-ray not scattered in patient (primary) = good event → retained
  𝛾-ray Compton scattered in patient (scattered)→ lost spatial information = bad event → rejected
Amplitude of PMT output signal proportional to energy of gamma ray 𝐸𝛾
  Amplitude of PMT output signal proportional to number of optical photons arriving from scintillator
  Number of optical photons produced in scintillator proportional to 𝐸𝛾
  Discriminating on the basis of amplitude of PMT output signal = discriminating on basis of 𝐸𝛾
Pulse height analyzer takes summed electrical signals from the PMT array and applies thresholds in two steps
  PMT output signals are sent to multiple-channel analyzer → pulse-height histogram
  Two thresholds, upper and lower, are applied to pulse-height histogram to select acceptable 𝛾-rays

## Slide 21: Pulse Height Analyzer (PHA)
Multiple channel analyzer (MCA), in which term ‘channel’ refers to specific energy range, digitizes signal within each range then produces pulse-height spectrum
  Pulse height analyzer allows spectroscopic imaging = selects for display only events falling within particular energy range
  Number of channels in MCA can be more than 1000, allowing essentially complete energy spectrum to be produced
  After digitization, upper and lower threshold values for accepting -rays are applied
Energy Resolution of system is defined as full-width-half-maximum (FWHM) of photopeak
Typically, about 14 keV (or 10%) for most gamma cameras without patient present
In clinical scan with patient in place, threshold level for accepting ‘photopeak’ is set to slightly larger value, typically 20%

## Slide 22: Instrument Dead Time
If an injected dose of radiotracer is very large, the total number of -rays striking the scintillation crystal can exceed the recording capabilities of the system
  Particularly true at the beginning of a clinical scan when radioactivity is at its highest level
  Limitation is due to the finite recovery and reset times required for various electronic circuits in the gamma camera
Two types of behaviour exhibited by the system components, termed paralyzable and nonparalyzable

## Slide 23: Instrument Dead Time
Paralyzable (extendable) behavior is a phenomenon in which a component of the system cannot respond to a new event until a fixed time after the previous one, irrespective of whether that component is already in a non-responsive state.
  Each time a -ray strikes the scintillation crystal it produces an excited electronic state, which decays in 230 ns to release a number of photons.
  If another -ray strikes same spot in the crystal before the excited state decays, then it will take a further 230 ns to decay, and only one set of photons will be produced even though two -rays have struck the crystal
  Very high rate of -rays striking detector can result in a considerably elongated dead-time and so the number of recorded events can actually decrease

## Slide 24: Instrument Dead Time
Non-paralyzable (non-extendable) component cannot respond for a set time, irrespective of the level of radioactivity
  For example, Anger logic circuit and PHA take a certain time to process a given electrical input signal, and during this time any further events are simply not recorded
  This time is fixed for a given system, and is not elongated if further scintillation events occur
Overall dead-time  is determined by non-paralyzable part
N: true count rate (number of scintillations per second), n: measured count rate,
maximum measurable count rate is given by 1/

## Slide 25: Instrument Dead Time Correction
Corrections for the dead-time are based upon calibrations using phantoms of known radioactivity, and also independent electronic measurements of the dead time of individual components of the system
Example: Suppose that the true count rate in a gamma camera is N=10,000 per second, but the measured rate is only n=8,000 per second. What is the dead time of the system?
Solution:

## Slide 26: Practical Planar Scintigraphy

## Slide 27: Single Photon Emission Computed Tomography (SPECT)
Injected radioactive tracer
Moving Gamma Camera
Tomographic scans can be used to reconstruct image

## Slide 28: Positron Emission Tomography (PET)
PET is a tomographic technique which also uses radiotracers
  Radiotracers used in PET emit positrons which, after annihilation with an electron in tissue, result in the formation of two -rays with energies of 511 keV each and with trajectories 180 apart
  PET has between 100-1000 times higher SNR and significantly better spatial resolution than SPECT

## Slide 29: Recommended Textbook
Introduction to Medical Imaging: Physics, Engineering and Clinical Applications, by Nadine Barrie Smith and Andrew Webb, Cambridge University Press, 2010.
  Relevant parts of Chapter 3