## Slide 1: Ultrasound Imaging
Prof. Yasser Mostafa Kadah – www.k-space.org
EE 472

## Slide 2: Pulse Echo Ultrasound Imaging
Echo Ranging
  Sending pulse and measuring time until reflected echo is received to determine location of reflecting surface
Time
Interface 1
Medium 1
Medium 2
Io
RA Io
(1-RA) Io

## Slide 3: A-Mode (Amplitude)
Sends ultrasound pulse and presents received echo as 1D plot
  Applications: opthalmology
Envelope Detection

## Slide 4: B-Mode (Brightness)
Transmits ultrasound pulse to a particular direction and reflected echoes are received and displayed as intensity values of one line
  Repeat in different directions to acquire multiple lines to form an image

## Slide 5: M-Mode (Motion)
One line in B-mode image is selected by doctor and scanned repeatedly with time – output is displayed as moving line trace
  Application: study of heart valves

## Slide 6: Doppler Mode
Spectral Doppler Mode (B&W Doppler)
Color Doppler Mode (Color Flow Mapping or CFM)
Power Doppler (Perfusion)
B&W Doppler
Color Doppler
Power Doppler

## Slide 7: 3D and 4D Imaging Modes
Volume Scanning
  Freehand or mechanical scanning
2D Imaging
3D Imaging
4D Imaging
Real-Time Slice
Static Volume
Real-Time Volume

## Slide 8: Ultrasound Physics
Sound waves used to form medical images are longitudinal waves, which propagate (travel) only through a physical medium (usually tissue or liquid)
  Characterized by frequency, wavelength, speed and phase

## Slide 9: Ultrasound Physics
Medical ultrasound frequencies used in the range 2–15 MHz
  Higher frequencies are now utilized for special applications
  Resolution proportional to wavelength
Acoustic impedance
  p is the local pressure and v is the local particle velocity.
  Analogous to electrical impedance (or resistance R )

## Slide 10: Ultrasound Physics
Reflection: Large Interfaces

## Slide 11: Ultrasound Physics
Scattering: Small Interfaces (size less than wavelength)
Two important aspects of scattering:
  Ultrasonic power scattered back is small compared to reflections
  Beam angle-independent appearance in the image unlike reflections
Diffuse Reflection: Rough Surfaces

## Slide 12: Ultrasound Physics
Refraction: Snell’s law
Attenuation: gradual loss of beam energy
  Depends on both distance and frequency

## Slide 13: Ultrasound Physics
Interference and diffraction
  Constructive/Destructive interference

## Slide 14: Ultrasound Physics
Focusing: narrower ultrasound beam

## Slide 15: Ultrasound Physics
Ultrasound pulse
Harmonic Imaging

## Slide 16: Ultrasound Physics
Acoustic pressure and intensities within ultrasound beam

## Slide 17: Ultrasound Physics
Acoustic pressure and intensities within ultrasound beam
  ISPTP
  ISPTA
  ISATA

## Slide 18: Ultrasound Transducer
Device that converts electrical transmission pulses into ultrasonic pulses and, conversely, ultrasonic echo pulses into electrical echo signals
  Piezoelectric material

## Slide 19: Ultrasound Transducer
Quarter-wavelength matching layer
Bandwidth for multi-frequency transducers

## Slide 20: Mechanical vs. Electronic Scanning

## Slide 21: Ultrasound Electronic Array Transducer
Linear- and convex-array transducers

## Slide 22: Electronic Scanning in Array Transducers
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32
Line 1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 | Line 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  | Line 3 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  | Line 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  | Line 5 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  | Line 6 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  | Line 7 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  | Line 8 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  | Line 9 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  | Line 10 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  | Line 11 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  | Line 12 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  |  | Line 13 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  |  |  | Line 14 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  |  |  |  | Line 15 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Line 16 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Line 17 |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Line 18 |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Line 19 |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Line 20 |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Line 21 |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Line 22 |  |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Line 23 |  |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Line 24 |  |  |  |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Line 25 |  |  |  |  |  |  | 
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32
Line 1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 3 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 5 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 6 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 7 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 8 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 9 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 10 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 11 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 12 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 13 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 14 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 15 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 16 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 17 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 18 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 19 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 20 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 21 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 22 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 23 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 24 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Line 25 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Linear Array
Phased Array

## Slide 23: Electronic Array Focusing: Beamforming
Transmission Focusing
Reception focusing
  Delay-Sum beamforming

## Slide 24: Dynamic Active Elements and Apodization
Use smaller number of elements for small depth
Use smaller element weights for outside elements to improve beam shape

## Slide 25: Grating lobes
Make sure to eliminate grating lobes by designing element spacing in array transducers to be half the ultrasound wavelength or less

## Slide 26: Dynamic Reception Focusing
Change focal point dynamically to follow received depth

## Slide 27: Multiple Transmission zones
Combine images with different focal points to improve overall beam shape
  Cost: lower frame rate

## Slide 28: Elevation Direction: Image Slide Thickness
Focusing is needed in the elevation direction as well
  Mechanical: Lens
  Electronic: 1.5D or 2D arrays

## Slide 29: Electronic Beamformer Design
Function: Compensation of different path lengths for echoes coming from a focal point to different array elements to emulate function of acoustic lens
  Delay and Sum beamforming: echoes from focal point arrive at same time
  Number of channels determines quality (e.g., 32 vs. 64)
Software Processing (Computer)
Analog Beamformer (Old Technology)
Digital Beamformer (New Technology)

## Slide 30: Frequency
Ultrasound attenuation (loss) in tissues is around 0.5 dB/cm/MHz
Image resolution is proportional to wavelength of ultrasound
Multi-frequency transducers offer flexibility to trade-off depth and resolution
  Transducers with 100% bandwidth are now common

## Slide 31: Gain (Brightness)
Increases intensity of received echoes at all depths
  Too little gain: too dark image
  Too much gain: too bright image
Proper Gain
Too Much Gain
Too Little Gain

## Slide 32: Time-Gain Compensation (TGC)
Amplifier with gain that varies with depth (time)
  Compensates for ultrasound attenuation inside tissues

## Slide 33: Dynamic Range (Contrast)

## Slide 34: B-Mode Instrumentation
Image reconstruction: scan conversion and interpolation
Real-time display: frame every 1/25 s
Freeze: updating frame stops
Cine Loop: recording of real-time scan as a movie
Frame Averaging: moving average filter to improve SNR

## Slide 35: B-Mode Image Properties
Lateral Resolution
Thickness resolution
  Elevation

## Slide 36: B-Mode Image Properties
Axial resolution
  Half pulse length
Speckle
  Random yet stationary pattern

## Slide 37: B-Mode Image Properties
Frame time / Frame rate
  Time to scan a complete image
  Example: time to scan 1 cm= 2x1cm/c= 2 cm/(1540 m/s) = 13 s
  Then, frame time to scan a 20 cm depth with 128 lines=13 s x20 x128
  Frame rate = 1/ frame time = 30 frames/s
Smaller D
Smaller N
Higher Frame Rate

## Slide 38: Doppler Ultrasound
Doppler effect: Change in the observed frequency of the sound wave compared to the emitted frequency which occurs due to relative motion between observer and source
Up-Shift
Down-Shift
No Shift

## Slide 39: Doppler Ultrasound
Doppler Effect: Change in received frequency of ultrasound wave compared to transmitted frequency due to relative motion between receiver and transmitter
Up-Shift
Down-Shift
No Shift

## Slide 40: Continuous Wave (CW) Doppler
Only a small region for Doppler sensitivity
No range information
No limitation on maximum velocity and high velocity accuracy

## Slide 41: Pulsed-Wave (PW) Doppler
Range information is available and region is selectable by user
Limitations on maximum velocity and accuracy

## Slide 42: Color Doppler
Display mean blood velocity at each point encoded in color
  Red: flow toward transducer
  Blue: flow away from transducer

## Slide 43: Power Doppler
Display estimate of power of all shifted components encoded in shades of orange
  Assessment of tissue perfusion

## Slide 44: Doppler Ultrasound
Doppler Shift Equation

## Slide 45: CW Doppler Signal Processing
Transmitted
Received
Transmitted x Received
Low-pass Filtered Transmitted x Received

## Slide 46: Clutter Problem in Doppler Ultrasound
Clutter: signal from stationary tissues
  Low Doppler shift and much stronger signal
  Signal from stationary tissue and wall motion
  Critical step in Doppler processing

## Slide 47: PW Doppler Signal Processing
Sampled version of CW Doppler
  Aliasing may occur

## Slide 48: PW Doppler Example
Consider Doppler imaging of a vessel at depth d1= 10 cm. derive the maximum detectible velocity if the transmitted signal frequency was 5 MHz and Doppler angle was 45.
Time to collect one sample = PRI = 13s/cm x (10 cm)= 130s
Sampling frequency = PRF = 1/PRI= 7692 Sa/s = 2 fdmax
vmax=(1540 x 7692/2)/(2x5x106xcos(45))

## Slide 49: Doppler Ultrasound Terminologies
Aliasing
  Highest Doppler frequency shift that can be measured is equal to PRF/2
Angle dependence
  Estimated Doppler shift is dependent on cosine of the angle between the beam and the direction of motion
Clutter breakthrough
  Tissue motion giving rise to Doppler frequencies above wall thump or clutter filter may be displayed on spectral Doppler or color flow systems
Loss of low Doppler
  Blood velocities which give rise to low Doppler frequencies (as a result of low velocity or angle near to 90°) will not be displayed if value of Doppler frequency is below the level of wall thump or clutter filter

## Slide 50: Ultrasound Safety and Limitations
Ultrasound waves may have mechanical or thermal effects on human tissues
  To date, no clear evidence that diagnostic ultrasound imaging causes harm
  Only imaging modality allowed to scan fetus
Fundamental approach to safety is to use lowest output power and shortest scan time consistent with acquiring required diagnostic information
  “ALARA” principle (i.e. as low as reasonably achievable)
Limitations of ultrasound imaging include:
  Inability to image tissues masked by air or bone/stones
  Geometric inaccuracy due to beam shape variation with depth
  Problems in imaging obese patients
A: 250 lb Patient
B: 150 lb Patient

## Slide 51: Quality Assurance
Performance testing using ultrasound phantom
  Uniformity
  Depth of Penetration
  Beam Profile/ Focal Zone/ Lateral Response Width
  Vertical Distance Measurement
  Horizontal Distance Measurement
  Axial and Lateral Resolution
  Elevational Resolution
  Contrast Resolution
  Grayscale Contrast Sensitivity
  Dead Zone Assessment

## Slide 52: Ultrasound Imaging System Components
Ultrasound Transducer
Amplifier (Gain)
Transmit Beamformer
Transmit/Receive Switch
Receive Beamformer
Preprocessing
Control
User Interface
Display
Pulser

## Slide 53: External Look
Mobile
Portable
Hand-Held

## Slide 54: Transducer

## Slide 55: Zero Insertion Force (ZIF) Connector
Easier to use for high pin count connectors to avoid damaging connector pins
  Common to have 156 connector pins or more

## Slide 56: User Interface: Keyboard and Display
Select, control, and show status
  Transducer
  Mode
  Frequency
  Depth
  Power
  Gain
  TGC
  Trackball
  Alphanumeric keyboard
  Freeze

## Slide 57: Covered Material
Diagnostic Ultrasound Physics and Equipment, Chapters 2, 3, 4, 5 and 7