## Slide 1: X-Ray Imaging
Prof. Yasser Mostafa Kadah
EE 472

## Slide 2: X-Ray Production

## Slide 3: X-Ray Production
Bremsstrahlung x-rays are produced when a projectile electron is slowed by the nuclear field of a target atom nucleus
  In the diagnostic range, most x-rays are bremsstrahlung x-rays
Characteristic x-rays are emitted when an outer-shell electron fills an inner-shell void
  This type of x-radiation is called characteristic because it is characteristic of the target element
  Only the K-characteristic x-rays of tungsten are useful for imaging
Approximately 99% of the kinetic energy of projectile electrons is converted to heat (Anode heat)

## Slide 4: Quantity and Quality of X-ray Beam
General shape of an emission spectrum is always the same, but its relative position along the energy axis can change
  The farther to the right a spectrum is, the higher the effective energy or quality of the x-ray beam
  The larger the area under the curve, the higher is the x-ray intensity or quantity

## Slide 5: Effect of mA and mAs
A change in mA or mAs results in a proportional change in the amplitude of the x-ray emission spectrum at all energies

## Slide 6: Effect of kVp
As kVp is raised, area under curve increases by approximating the square of the factor by which kVp was increased
  Accordingly, x-ray quantity increases with the square of this factor
Change in kVp affects both amplitude and position of x-ray emission spectrum
  In diagnostic range, 15% increase in kVp is equivalent to doubling mAs

## Slide 7: Effect of Added Filtration
Adding filtration to the useful x-ray beam reduces x-ray beam intensity while increasing the average energy
  The result of added filtration is an increase in the average energy of the x-ray beam with an accompanying reduction in x-ray quantity

## Slide 8: Effect of Target Material
The atomic number of the target affects both the number (quantity) and the effective energy (quality) of x-rays
  As the atomic number of the target material increases, the efficiency of the production of bremsstrahlung radiation increases, and high-energy x-rays increase in number to a greater extent than low-energy x-rays.

## Slide 9: Effect of Voltage Waveform
There are five voltage waveforms: half-wave–rectified, full-wave–rectified, three-phase/six-pulse, three-phase/12-pulse, and high-frequency waveforms
Both quantity and quality decrease by ripple
  Because of reduced ripple, operation with three-phase power or high frequency is equivalent to an approximate 12% Increase in kVp, or almost a doubling of mAs over single phase power.

## Slide 10: Factors Affecting X-Ray Quantity

## Slide 11: Factors Affecting X-Ray Quality

## Slide 12: Half-Value Layer (HVL)
In radiography, quality of x-rays is measured by the HVL
  Diagnostic x-ray usually has HVL 3 to 5 mm Al or 3 to 6 cm of soft tissue
Although x-rays are attenuated exponentially, high-energy x-rays are more penetrating than low-energy x-rays
  100-keV x-rays are attenuated at rate of 3%/cm of soft tissue
  10-keV x-rays are attenuated at 15%/cm of soft tissue

## Slide 13: X-Ray Interaction with Matter
Coherent scattering (energy < 10 keV)
Compton scattering
Photoelectric effect
Pair production (energy > 1.02 MeV)
Relevant to x-ray imaging

## Slide 14: Compton (Incoherent) Scattering
+

## Slide 15: Photoelectric Effect

## Slide 16: Photoelectric Effect

## Slide 17: Differential Absorption

## Slide 18: X-Ray Exponential Attenuation
Total reduction in number of x-rays remaining in x-ray beam after penetration through a given thickness of tissue is called attenuation
When broad beam of x-rays is incident on any tissue, some x-rays are absorbed, and some are scattered
Attenuation is the combined effect of both absorption and scattering
: Attenuation coefficient
: Path length
Tissue

## Slide 19: Radiologic Units
Air Kerma (Kinetic Energy Released in Matter) (Gya)
  Kinetic energy transferred from photons to electrons during ionization and excitation measured in J/kg where 1 J/kg = 1 gray (Gya)
Absorbed Dose (Gyt )
  Radiation energy absorbed in tissue per unit mass with units of J/kg or Gyt (gray) which depends on tissue type
Sievert (Sv): quantity of radiation received by radiation workers and populations
Becquerel (Bq): quantity of radioactive material, not the radiation emitted by that material
  Radioactivity and the becquerel have nothing to do with x-rays

## Slide 20: Radiologic Units

## Slide 21: X-Ray Tube
External structures
  Support structure
  Protective housing
  Glass or metal enclosure
Internal structures
  Anode and cathode

## Slide 22: X-Ray Tube Support Structure
X-ray tube and housing assembly are quite heavy
  Require support mechanism so radiologic technologist can position them
  Mainly ceiling, floor or C-arm support systems

## Slide 23: Protective Housing
When x-rays are produced, they are emitted isotropically
  That is, with equal intensity in all directions
Only x-rays emitted through window are called useful beam
  X-rays that escape through protective housing: leakage radiation
  Leakage radiation contributes nothing to diagnostic information and result in unnecessary exposure of patient and radiologic technologist
Protective housing guards against excessive radiation exposure and electric shock
  Also mechanically protects x-ray tube

## Slide 24: Metal or Glass Enclosure
X-ray tube is an electronic vacuum tube with components contained within a glass or metal enclosure
  vacuum allows for more efficient x-ray production and longer tube life
As glass enclosure tube ages, some tungsten vaporizes and coats the inside of glass enclosure
  Alter electrical properties of the tube, allowing tube current to stray and interact with the glass enclosure resulting in arcing and tube failure
  Most common cause of tube failure
Metal enclosures maintain constant electric potential between electrons of tube current and enclosure
  Longer life and less likely to fail
  Virtually all high-capacity x-ray tubes now use metal enclosures

## Slide 25: Cathode
Cathode is the negative side of the x-ray tube
It has two primary parts, a filament and a focusing cup
  Dual-filament cathode allows two focal spots (e.g., 0.5 and 1.5 mm)
  Focusing cup is a metal shroud that surrounds filament
  Tube current is adjusted by controlling filament current

## Slide 26: Anode
Anode is the positive side of the x-ray tube
Two types: stationary (dental) and rotating (general purpose)
  Higher tube currents and shorter exposure times are possible with rotating anode because of their better heat dissipation
Three functions in an x-ray tube:
  Electrical conductor that receives electrons emitted by cathode and conducts them through the tube to the connecting cables and back to the high-voltage generator
  Mechanical support for the target
  Thermal dissipation

## Slide 27: Target
The target is area of anode struck by electrons from cathode

## Slide 28: Focal Spot
Focal spot is the area of target from which x-rays are emitted
  The smaller the focal spot, the better the spatial resolution of the image
  Unfortunately, as the size of focal spot decreases, heating of target is concentrated onto a smaller area (limiting factor to focal spot size(
Line-focus principle: angling target makes effective area of the target much smaller than actual area of electron interaction

## Slide 29: Radiographic Image Quality
Definition: fidelity with which anatomical structure being examined is rendered on radiograph
  Spatial resolution: ability to image small objects
  Contrast resolution: ability to distinguish anatomical structures
  Radiographic noise: random fluctuation in intensity of image
    Film graininess, structure mottle, quantum mottle, and scatter radiation

## Slide 30: Geometric Factors: Magnification

## Slide 31: Geometric Factors: Distortion
Unequal magnification of different portions of the same object is called shape distortion
  Distortion depends on object thickness, position, and shape
  Thick objects are more distorted than thin objects
  If object plane and image plane are not parallel, distortion occurs

## Slide 32: Geometric Factors: Focal-Spot Blur
Focal-spot blur is caused by effective size of focal spot
  The most important factor for determining spatial resolution
  Smaller on anode side than cathode side of the image (Heel effect)

## Slide 33: Subject Factors
kVp is the most important influence on subject contrast

## Slide 34: Control of Scatter Radiation
Reduced image contrast results from scattered x-rays
  Restricting x-ray beam (collimation) reduces scattering

## Slide 35: Beam Restricting Devices
Collimation reduces patient radiation dose and improves contrast resolution

## Slide 36: Radiographic Grids
Effective device for reducing level of scatter radiation that reaches image receptor
  The principal function of a grid is to improve image contrast

## Slide 37: Radiographic Grids
High-ratio and high-frequency grids increase patient radiation dose
When grid is used, radiographic technique must be increased to produce same image receptor signal by a factor called Bucky (Grid) factor (B)
  As Bucky factor increases, radiographic technique and patient dose increases
  The higher the grid ratio, the higher is the Bucky factor
  The Bucky factor increases with increasing kVp

## Slide 38: Radiographic Grids
Grid Cutoff: undesirable absorption of primary x-rays by grid
  Greater Attenuation of primary x-rays near edges of image receptor

## Slide 39: Radiographic Grids Types
Parallel, Crossed and Focused
Moving Grid (Bucky): reciprocating and oscillating
  (-) Require a bulky mechanism that is subject to failure
  (-) Distance between patient and the image receptor is increased
  (-) Moving grids can introduce motion into cassette-holding device
  Advantages of moving grids far outweigh disadvantages

## Slide 40: Computed Radiography (CR)
Filmless radiology using special imaging plates
  Photostimulable luminescence (PSL)

## Slide 41: Computed Radiography (CR)
Computed
Radiography
Screen-Film Radiography
Proper radiographic technique and exposure are essential
Radiographic technique is not so critical

## Slide 42: Digital Fluoroscopy (DF)
Fluoroscopy: real-time dynamic viewing of anatomic structures
  Advantages of DF: speed of image acquisition and postprocessing to enhance image contrast

## Slide 43: Interventional Radiology
Performing surgical procedures under guidance from radiographic equipment

## Slide 44: Digital Mammography
Radiographic examination of the breast
Digital Mammography spatial resolution limited by pixel size
  Superior contrast resolution principally because of postprocessing

## Slide 45: Operating Console
Allows radiologist to control x-ray tube current and voltage so that useful x-ray beam is of proper quantity and quality
  Radiation quantity refers to number of x-rays or intensity of x-ray beam
  Radiation quality refers to penetrability of x-ray beam and is expressed in kilovolt peak (kVp) or, more precisely, half-value layer (HVL)

## Slide 46: Autotransformer
Power supplied to x-ray imaging system is delivered first to autotransformer where it provides controlled but variable voltage to high-voltage transformer
  It is much safer and easier to control a low voltage and then increase it than to increase a low voltage to the kilovolt level and then control its magnitude

## Slide 47: Adjustment of Kilovolt Peak (kVp)
kVp determines the quality of the x-ray beam
Appropriate autotransformer connections can be selected with an adjustment knob, a push button, or a touch screen
  This low voltage from autotransformer becomes the input to high-voltage step-up transformer that increases voltage to chosen kilovolt peak
Note: kVp meter placed across output terminals
of autotransformer actually reads voltage, not kVp.
It registers kilovolts because of the known multiplication factor of high voltage transformer

## Slide 48: Control of Milliamperage (mA)
The x-ray tube current, crossing from cathode to anode, measured in milliamperes (mA)
  Number of electrons emitted by filament is determined by filament temperature (controlled in turn by filament current)
  Thermionic emission is the release of electrons from a heated filament
Space Charge Effect: As the kVp is raised, anode becomes more attractive to electrons that
would not have enough energy to leave the filament. Hence, this effectively increases mA with kVp and hence should be corrected for by special circuit

## Slide 49: Exposure Timer
Most exposure timers are electronic, controlled by microprocessor
  Allow wide range of time intervals to be selected and are accurate to intervals as small as 1 ms
Special kind of electronic timer, called an mAs timer, monitors product of mA and exposure time and terminates exposure when desired mAs value is reached
  Because the mAs timer must monitor actual tube current, it is located on the secondary side of the high-voltage transformer

## Slide 50: Automatic Exposure Control (AEC)
AEC is a device that measures quantity of radiation that reaches image receptor and automatically terminates exposure when image receptor has received required radiation intensity

## Slide 51: High-Voltage Generator
Function: increases output voltage from autotransformer to kVp necessary for x-ray production
High-voltage generator contains three primary parts: high-voltage transformer, filament transformer, and rectifiers
Note: Although some heat is generated in the high-voltage section and is conducted to oil,
the oil is used primarily for electrical insulation

## Slide 52: High-Voltage Transformer
High voltage transformer is a step-up transformer
  Turns ratio of is usually between 500:1 and 1000:1

## Slide 53: High-Voltage Rectification
Rectification is the process of converting AC to DC
  Rectification is accomplished with diodes
Transformers operate AC while x-ray tubes need DC
  X-rays are produced by acceleration of electrons from cathode to anode and cannot be produced by electrons flowing in reverse

## Slide 54: Single-Phase vs. Three-Phase
Three-phase power is a more efficient way to produce x-rays than is single-phase power
  With three-phase power, voltage applied across the x-ray tube is nearly constant, never dropping to zero during exposure.

## Slide 55: High-Frequency Generator
High-frequency generators produce nearly constant potential voltage waveform, improving image quality
Rectified power at 60 Hz is inverted to a higher frequency, from 500 to 25,000 Hz, then transformed to high voltage
  Advantage: much smaller size than 60-Hz high-voltage generators

## Slide 56: Voltage Ripple Comparison
Less voltage ripple results in greater radiation quantity and quality

## Slide 57: Power Rating
Transformers and high-voltage generators usually are identified by their power rating in kilowatts (kW)
  Power (W) = Current (A) × Potential (V)
For specifying high-voltage generators, the industry standard is to use the maximum tube current (mA) possible at 100 kVp for an exposure of 100 ms
  This generally results in the maximum available power
Use RMS voltage factor to account for voltage ripples
  0.7 of peak in single phase generators
  Close enough to 1 in three-phase and high-frequency generators

## Slide 58: X-Ray Circuit

## Slide 59: Cardinal Principles for Radiation Protection
Simplified rules designed to ensure safety in radiation areas for occupational workers

## Slide 60: Cardinal Principles for Radiation Protection
Minimize Time
  Dose is directly related to duration of radiation exposure
  Exposure = Exposure rate × Exposure time
Maximize Distance
  As distance between source of radiation and person increases, radiation exposure decreases rapidly by inverse square law
  If distance from source exceeds 5 times source diameter, it can be treated as point source (assume true and apply inverse square law)
Use Shielding
  Positioning shielding between radiation source and exposed persons greatly reduces level of radiation exposure
  Shielding used in diagnostic radiology usually consists of lead, although conventional building materials also are used

## Slide 61: Shielding
Estimate dose reduction using half-value layer (HVL) or tenth-value layer (TVL) of barrier material (1 TVL = 3.3 HVL)
Protective apparel
  Protective aprons usually contain 0.5 mm Pb (2 HVL – reduction to 25%).
  Actual measurements show reduction to approximately 10%

## Slide 62: Effective Dose
Effective dose is the equivalent whole-body dose
  When only part of body is exposed, as in medical x-ray imaging, risk is proportional to effective dose (E)
  Equivalent whole-body dose is the weighted average of the radiation dose to various organs and tissues

## Slide 63: Patient and Occupational Effective Dose

## Slide 64: Covered Material and Suggest Problems
Stewart C. Bushong, Radiologic Science for Technologists: Physics, Biology, and Protection, 11th ed., Mosby, 2017. (ISBN 978-0-323-35377-9)
  Chapters 1, 5, 6, 7, 8, 9, 10, 11, 35