## Slide 1: Human Factors Engineering: Design of Medical Devices Signs, Symbols, and Markings
Prof. Yasser Mostafa Kadah – www.k-space.org

## Slide 2: Recommended Reference
ANSI/AAMI HE75: 2009

## Slide 3: Introduction
Proper labeling can have an enormous impact on how a user interacts with a device by increasing intuitiveness, minimizing required training, and ensuring accurate performance
Properly designed labeling presents users with instructions, identifications, markings, symbols, colors, and commands that they already know or can easily intuit
  Labeling of medical devices and associated accessories, shipping containers, and wrappings
Labeling should be the absolute last resort for protecting users from potential risks

## Slide 4: General Considerations
All equipment elements that are not intuitively obvious, that have some usable functions, or that require identification for other reasons should be labeled or marked
Markings should be positioned so that they are clearly associated with the correct equipment feature and not obscured by hand positions or equipment components
Appropriate markings can enhance the identification of both individual elements and their functional relationships
Labels should be visible at typical viewing distances and angles

## Slide 5: General Considerations
Labels should be resistant to wear and tear
Ambiguous symbols, codes, or terminology should be avoided
Designing for legibility requires careful analysis of ambient illumination in typical use environments
Consistency of placement, terminology, and coding is critical
All markings should be tested with typical users. Users can be clinicians, caregivers, patients, or maintenance personnel and can vary by age, disability, and other characteristics

## Slide 6: Label Design Model

## Slide 7: Design Guidelines
Labels for equipment identification
  At a minimum, provide the name or trade name of the manufacturer, a distinctive catalog or model number, and the electrical rating
  Positioning of nonfunctional markings (e.g., logo) should not distract from or compromise the clear presentation of other information

## Slide 8: Design Guidelines
Hazard labels
  Hazards categorized as (a) death or irreversible injury, (b) reversible injury, and (c) minor injury or discomfort
  Used to alert personnel to possible hazards that could be encountered during the use, handling, storage, or repair of a device

## Slide 9: Design Guidelines
Electrical receptacle and connector labels
  Receptacles and connectors should be marked with their power rating and the intended function or connecting cable
  Convenience receptacles should be labeled with the maximum allowable load in amperes or watts
Fuse and circuit-breaker labels
  Permanently marked adjacent to the fuse holder
  Legible in the ambient illumination anticipated

## Slide 10: Design Guidelines
Labels on controls, keyboards, and keypads
  Intended to facilitate identifying, locating, and functionally grouping user display and control components.
  Markings of control positions are also important to users when they adjust mechanical controls and read mechanical displays
Positioning and mounting of labels
  Placed near above panel elements, panel displays, or flow lines that users grasp, press, follow, or otherwise handle
  Positioned for visibility and readability
  from typical viewing angles
  Separated by sufficient space

## Slide 11: Design Guidelines
Label orientation
  Labels should be oriented horizontally so that they can be read quickly
Indications of functional relationships
  “Location aids” such as demarcation, color coding, background shading, mimics, and flashing lights can be used to indicate the positions of, and relationships between, functionally related controls, displays, and other features
  Redundant if necessary to ensure reliability

## Slide 12: Design Guidelines
Permanence and durability of labels
  Markings should be permanent and should remain legible throughout the intended life of the equipment under anticipated use and maintenance conditions
  Durability is of special concern for controls like keypads and pushbuttons because constant use can lead to wear
  Etched molding

## Slide 13: Specific Design Guidance
Consistency
  Wording, symbology, coding, and all other aspects of labeling should be consistent within devices and systems.
Label content
  Clarity, consistency, and brevity of label information are critical

## Slide 14: Specific Design Guidance
Use of standard symbols
  ANSI/AAMI/IEC TIR 60878

## Slide 15: Specific Design Guidance
Legibility
  Contrast, lettering, character height (refer to Ch. 6)
  Assessed under worst-case viewing conditions (i.e., max intended viewing distance and angle, worst-case lighting, and min exposure time)
Coding
  Use of shapes, colors, symbols, or other attributes that have established meanings or associations for a given user population.
  Facilitates quick, accurate identification of controls, displays, connector ports, and other features.
Application of color to mimics and flow lines
  Different colors with sufficient contrast
  Lines depicting the flow of the same contents (e.g., blood, oxygen) should have the same color throughout the system

## Slide 16: Specific Design Guidance
Package labels
  Should provide the product and model names and the name and business address of the manufacturer, packer, or distributor.
  Should also describe shipping and storage requirements.

## Slide 17: Further Reading and Assignments
Chapter 10 of Recommended Reference
Solved MCQ exams on web site (Part of Medical Equipment courses)