## Slide 1: Mechanical Ventilator: The Basics
Professor Yasser Mostafa Kadah

## Slide 2: Mechanical Ventilator
A ventilator is a life-sustaining device that supports or replaces spontaneous breathing of the patient

## Slide 3: Spontaneous Breathing Cycle
A breath is one cycle of positive flow (inspiration) and negative flow (expiration) defined in terms of the flow versus time curve

## Slide 4: Gas Exchange: Diffusion
Gas is exchanged between alveoli and blood by diffusion
Speed of diffusion is determined by:
  Partial pressure gradient (+)
  Thickness of barrier (-)
In cases when fluid accumulates inside alveoli, hyperoxic (more O2) is needed to compensate for increased barrier thickness

## Slide 5: Equations of Motion: Electrical Analogy
Mechanics Electricity
Pressure  Voltage
Flow  Current
Volume  Charge
Resistance  Resistance Compliance  Capacitance
Valve  Switch

## Slide 6: Practical Effects of Compliance and Resistance
A: Filling of normal lung unit
B: Low-compliance unit: fills quickly but with less air
C: Increased resistance: fills slowly
Inspiration
Expiration

## Slide 7: Functional Block Diagram

## Slide 8: Oxygen Sources
Medical gases supply
Oxygen cylinder
Oxygen concentrator

## Slide 9: Air/Oxygen Blender
Gas mixer allow the user to vary oxygen concentration of inspiratory gas or fractional inspiratory oxygen (FIO2) between 21% and 100% by volume
  Mechanical gas mixers
  Electronically-controlled gas mixer integrated in ventilator
Gas mixers are usually responsible for ensuring that breathing gas to be supplied is prepared and delivered in required quantity and rate

## Slide 10: Pressure/Flow Generator
Pressure/flow generator is responsible for delivering mixed gas prepared by gas mixer according to selected ventilation parameters
Flow generator is a controlled valve whose output provides defined gas flow with output pressure is not specified
Pressure generator behaves similar to compressor, whose output provides defined pressure with unspecified gas flow

## Slide 11: Breathing System
Breathing system forms interface between patient and ventilator
Clinical ventilators are usually connected to patient via inspiratory and expiratory hose (dual-hose circuit)
Gas flow delivered through inspiratory port passes through breathing gas humidifier before entering patient’s lungs
  To make it adapted to climatic conditions in patient’s lungs
After inspiratory phase, patient exhales through expiratory valve
  Breathing systems of ventilators are referred to as non-rebreathing circuits

## Slide 12: Gas Humidifier
Humidifiers are used to warm and humidify inspiratory gas.
  Dry and cool supply gas dry out the patient’s airways with risk of causing irreversible damage to ciliated epithelium
Active gas humidifiers
  Located in the inspiratory limb and use electrical energy to heat a water bath
  When cold, dry gas passes over the water surface it absorbs water molecules and is thus warmed and humidified
  Examples: pass-over humidifiers and bubble-through humidifiers

## Slide 13: Gas Humidifier
Passive breathing gas humidifiers or heat and moisture exchangers (HMEs)
  Designed to buffer significant fraction of moisture and heat expired by patient
  Retained moisture is then used to condition inspired gas passing through HME during next inspiration

## Slide 14: Expiratory (Exhalation) Valve
Allows expired air to move out of breathing circuit
If valve is not opened completely during expiration, positive end-expiratory pressure (PEEP) is created in lungs
  PEEP is therapeutically important as it increases gas exchange surface of lungs
  Adequate PEEP can also prevent collapse of individual alveolar areas
If expiratory valve is controlled during inspiratory phase, it can compensate for undesired pressure rises in breathing system
  Caused, for example, by patient coughing

## Slide 15: Patient Interface
Noninvasive Ventilation (NIV) vs. Intubation

## Slide 16: Infection Control: HEPA Filtration
Filtration is the passage of a liquid or gas through a filter with pores small enough to retain microbes
  Membrane filters composed of cellulose esters are commonly used to filter out bacteria, viruses
Microbes can be removed from air by high-efficiency particulate air (HEPA) filters
  Remove almost all microorganisms larger than about 0.3 m in diameter

## Slide 17: Trigger, Limit, Cycle, and Baseline Variables
Trigger variable is one that is measured and used to start inspiration
Limit variable is one that can reach and maintain a preset value before inspiration ends (i.e., does not end respiration)
Cycle variable is one that is measured and used to end inspiration
Baseline variable is the parameter controlled during expiration
  Pressure control is most practical and used in all modern ventilators

## Slide 18: Modes of Mechanical Ventilator
“Mode” is defined as a predetermined pattern of interaction between a ventilator and a patient
  There are over 100 names for modes of ventilation on commercially available mechanical ventilators
  Neither the manufacturing community nor the medical community adhere to a standard taxonomy for modes
Volume (VC)
Pressure (PC)
CMV
IMV
CSV
Set-point
Adaptive
Servo
Intelligent
Ventilator Mode
+
+
=
Breath Control Variable
Breath Sequence
Targeting Scheme

## Slide 19: Modes of Mechanical Ventilator

## Slide 20: Beginning of Inspiration: Trigger Variable
Time triggering
Patient triggering
  Pressure
  Flow
  Volume

## Slide 21: During Inspiration: Control and Limit Variables
Ventilator can control one variable
  Pressure, volume, flow, or time
Limit variable is maximum value variable can attain during inspiration
  Does not end inspiratory phase
  Ventilator can have multiple limit variables
Control
Control

## Slide 22: Termination of Inspiration: Cycle Variable
Volume-Cycled Ventilation
Time-Cycled Ventilation
Flow-Cycled Ventilation
Pressure-Cycled Ventilation

## Slide 23: Targeting Scheme

## Slide 24: Example Ventilator Modes
CPAP
Spontaneous
Pressure-Support (PS)
Continuous Mandatory Ventilation (CMV)
Synchronized Intermittent Mandatory Ventilation (SIMV)

## Slide 25: CPAP, PEEP and BiPAP
Continuous Positive Airway Pressure (CPAP)
Positive End-Expiratory Pressure (PEEP or EPAP)
BiLevel Positive Airway Pressure (BiPAP)
BiPAP
CPAP/PEEP + Spontaneous
PEEP + CMV
CPAP/PEEP + IMV

## Slide 26: Typical Ventilator User Interface
Mode
IPAP/EPAP
Rate
FIO2
Tidal volume
Pressure, flow, volume
Alarm messages

## Slide 27: Recommended Further Reading
J.M. Cairo, Pilbeam's Mechanical Ventilation: Physiological and Clinical Applications, 5th Ed., Elsevier, 2012.
Martin J. Tobin, Principles And Practice of Mechanical Ventilation, 3rd Ed., McGraw Hill Medical, 2012.

## Slide 28: Thank You